## Slide 1: A first look at using the Vapor Sorption Analyzer in soils
Doug Cobos, Ph.D.
Decagon Devices and Washington State University

## Slide 2: Soil Moisture Characteristic Curve (SMCC)
Moisture release curve, water retention function, pF curve, moisture sorption isotherm
Relationship between water content and water potential (water activity, suction, pF, chi)

## Slide 3: Generating SMCC
Measuring water content is easy
  Gravimetric analysis (oven drying)
Measuring water potential is difficult
  No single instrument can make accurate measurements from wet to dry

## Slide 4: Wet end instruments (liquid equilibrium)
Pressure chamber
Tensiometers
Wind-Schindler
(HyProp)

## Slide 5: Dry end instruments (vapor equilibrium)
Chilled mirror dew-point
hygrometer
Thermocouple psychrometers
Relative humidity (hr) and water potential (Ψ) related by the Kelvin equation:
R is universal gas constant
Mw is molecular mass of water
T is temperature

## Slide 6: Dry end SMCC
Historically very difficult to obtain
  Campbell and Shiozawa (1992) cited over 70 times and data used many by many authors
  Introduction of WP4, WP4T, WP4C made dry end SMCC more accessible
Aquasorp/VSA instruments are the next step

## Slide 7: VSA (Vapor Sorption Analyzer)
Generates dry end SMCC
Fully automated
Drying and wetting (hysteresis loop)

## Slide 8: VSA limitation
Limitation: only works drier than -7 MPa (95% relative humidity)

## Slide 9: VSA (how it works)
Infrared Sensor
Mirror
Optical Sensor
Fan
Sample
Dry Air
Wet Air
Precision Balance
Dynamic mode
  Wet (~99% rh) or dry (~0%rh) air flows across sample
  Flow stops and water activity/water potential and mass (water content) measured
  Extreme resolution (> 200 points)
  < 48 hours for wetting and drying loop (often < 10 hours)

## Slide 10: VSA (how it works)
Infrared Sensor
Mirror
Optical Sensor
Fan
Sample
Dry Air
Wet Air
Precision Balance
Static mode
  Humidity of chamber controlled at pre-determined level
  Mass measured over time until no longer changing
  Less resolution
  Sorption kinetics
Sorption/desorption of other gases is possible

## Slide 11: Static vs. Dynamic
Chart: Water Activity
| 9.890000000000021E-2 | 0.19950000000000001 | 0.29960000000000003 | 0.40060000000000001 | 0.50029999999999997 | 0.60040000000000004 | 0.70140000000000002 | 0.80049999999999999 | 0.70050000000000001 | 0.60040000000000004 | 0.50019999999999998 | 0.39810000000000106 | 0.30050000000000004 | 0.2001 | 9.890000000000021E-2
Static Method | 3.2163 | 3.825 | 4.4261 | 5.3265 | 6.4181 | 7.306 | 8.3845 | 9.6553 | 8.812 | 7.9127 | 6.9791 | 6.1768 | 5.369 | 4.478 | 3.3847

## Slide 12: Static method (kinetics)

## Slide 13: Dynamic method

## Slide 14: What can you do with VSA data? Hysteresis

## Slide 15: Slope of semi-log plot
Logarithm of water potential vs. water content
  Plot becomes straight line in dry (VSA) region
  Slope of plot contains much useful information

## Slide 16: pF and Chi (C) Condon (2006), Orchiston, 1953
pF = log (Ψ) where Ψ has units of cm of water
  Common measure in Europe and in geotechnical engineering community
  Takes log of number with units (mathematical mistake)
  Increases with decreasing moisture (not intuitive)
C = -ln[-ln(aw)] where aw is water activity (relative humidity)
  Mathematically correct, decreases with decreasing moisture

## Slide 17: Slope of semi-log plot
Adsorption leg of semi-log plot extends through -1000 MPa at 0 water content
  pF 7.0, Chi = -2.00
  aw = 0.0006, relative humidity = 0.06%
Slopes highly variable by soil type
Chart: Chi -ln(-ln(aw))
| -2 | 1 | -2 | 1 | -2 | 1 | -2 | 1 | -2 | 1 | -2 | 1
 | 0 | 1.3
 | 0 | 2.4
 | 0 | 2.8
 | 0 | 5
 | 0 | 7.3
 | 0 | 11.3

## Slide 18: What can you do with VSA data? Specific Surface Area (SSA)
BET model (Brunauer et al. 1938, Likos and Lu 2002)
  Physically based model for isotherms (SMCCs)
Xm is water content when soil covered with monolayer of water
Water content
Water activity
offset
slope

## Slide 19: What can you do with VSA data? Specific Surface Area (SSA)
  Slope of Chi plot (Condon 2006)
SSA= f*S*a
S is slope of Chi plot (g/g)
a is monolayer coverage (3500 m2 g-1)
f is factor of 1.84

## Slide 20: What can you do with VSA data? Specific Surface Area (SSA)
Tuller and Or (2005)
  General scaling model for SSA from SWCC
  Mixed results in literature
Resurreccion (2011)
  Used slope of log water potential vs. water content function for SSA

## Slide 21: What can you do with VSA data? Specific Surface Area (SSA)
Which SSA measure is the best?
Still much research to be done!

## Slide 22: What can you do with VSA data? Adsorbed cation type
Smectite exchanged to achieve homoionic state
Fundamentally different isotherms

## Slide 23: What can you do with VSA data? CEC (cation exchange capacity)
“Intrinsic Isotherm” approach
  Water content normalized by dividing by CEC of soil
  All isotherms converge
  Maybe inverse relationship could be used to predict CEC?
  More research needed!

## Slide 24: What can you do with VSA data? Soil clay content (clay activity)
Resurreccion (2011)
  Used slope of log water potential vs. water content function for SSA
Decagon internal research
  Used slope of chi plot

## Slide 25: What can you do with VSA data? Gas movement in soil
Simulations of water vapor transport for pesticide volatilization (Chen et al., 2000)
Remediation of Volatile Organic Carbon compounds (Batterman et al., 1995)

## Slide 26: What can you do with VSA data? Swelling potential (expansiveness)
McKeen (1992) showed that the slope of pF vs. water content is related to soil swelling potential
Class | Slope | Expansion
I | > -6 | special case
II | -6 to -10 | high
III | -10 to -13 | medium
IV | -13 to -20 | low
V | < -20 | non-expansive

## Slide 27: What can you do with VSA data? Swelling potential (expansiveness)
Other possible swelling potential indicators
  Breadth and area of hysteresis loop?
  Monolayer coverage and heat of hydration from BET analysis?
  Equilibrium water content at specific relative humidity
  Sorption rate constants derived from step change in relative humidity (possible with new VSA)
  MUCH more research needs to be done!

## Slide 28: What can you do with VSA data? Unexplained results
Bentonite SMCCs (isotherms)
  Desorption curve independent of starting water activity
  Adsorption curve totally dependent on starting water activity

## Slide 29: Take-home points
Aquasorp/VSA data useful for
  Specific surface area
  CEC
  Clay activity
  Swelling potential
  Modeling of gas transport in soil
  Sorption kinetics (not even explored yet)
We have not even begun to understand the usefulness of dry end SMCC/isotherms

## Slide 30: References
Batterman, S.A., A. Kulshrethsa, and H.Y. Chang. 1995. Hydrocarbon vapour transport in low moisture soils . Environmental Science and Technology 29: 171-180
Brunauer, S., P. H. Emmett, and E. Teller. 1938. Adsorption of gases in multi-molecular layers. J. Am. Chem. Soc. 60:309-319.
Chen, D., D.E. Rolston, and P. Moldrup. 2000. Coupling diazinon volatilization and water evaporation in unsaturated soils: I. Water transport, Soil Science 165: 681-689
Condon, J. B. 2006. Surface Area and Prorsity Determination by Physisorption: Measurements and Theory. Elsevier, Amsterdam.
Likos, W. J. and N. Lu. 2002. Water vapor sorption behavior of smectite-kaolinite mixtures. Clays and Clay Minerals 50: 553-561.
Orchiston, H. D. 1953. Adsorption of water vapor: I. Soils at 25 C. Soil Sci. 73:453-465
Resurrecction, A. C., P. Moldrup, M. Tuller, T.P.A. Ferre, K. Kawamoto, T. Komatsu, L. W. de Jonge. Relationship between specific surface area and the dry end of the water retention curve for soils with varying clay and organic carbon contents. Water Resources Research 47, W06522
Tuller, M. and D. Or. 2005. Water films and scaling of soil characteristic curves at low water contents. Water Resources Research 41, W09403.

## Slide 31: Appendix: wet end measurements
Range can be extended further into wet end
  Start with saturated sample and dry down
  Can only get drying leg of hysteresis loop
    VSA can’t wet back up past about -10 MPa using vapor equilibration

## Slide 32: Appendix: wet end measurements

## Slide 33: Appendix: wet end measurements