## Slide 1: Decagon Instrument Assessment
Chris Chambers

## Slide 2: Basic Troubleshooting
Power?
Signal?
Accuracy?

## Slide 3: EM50 troubleshooting flow chart
Is the Status light blinking?
Replace batteries and press reset button
No
Attempt to connect to logger
Yes
Status light comes on?
No
Contact Decagon for an RMA
Connection made?
Yes
No
Correct Com port selected?
Device Driver installed?
Cable ok?
Error message?
Ports?
Sensors?
Communications?

## Slide 4: EM50G Telemetry
If you can connect to the logger, but cannot remotely download data from an EM50G:
  Run a telemetry test. Note any messages returned.
  If test failed, launch Ech2o Utility, go to Actions and select “List cellular carriers”. Is there any carrier compatible with your SIM plan?
  Contact support@decagon.com and provide:
    Device ID
    Approximate location of logger
    Error messages returned by telemetry test

## Slide 5: Sensor Troubleshooting
There are no known cases of our sensors drifting over time.
They usually either work or they do not.
#1 cause of sensor problems is cable damage

## Slide 6: Soil moisture Sensors
Sensor | Air reading | Water
EC5 | 323 | 1300
10HS | 470 | 1550
5TE/5TM* | 1 | 80
GS3* | 1 | ~75
GS1 | 1380 | 3150

## Slide 7: If the customer is using a non-Decagon logger
Check wiring!

## Slide 11: Firmware updates
I absolutely recommend regular firmware updates for:
  Loggers
  Instruments (LP-80, SC-1, ProCheck, etc.)
I do not recommend sensor updates unless there is a known problem that it may solve

## Slide 12: Other troubleshooting advice
Check the manual!
A correct date and time and location is critical for some instruments
  LP-80
Calibrations need to be done correctly and regularly
  LP-80
  SC-1

## Slide 13: Return Material Authorization (RMA)
support@decagon.com, ponczoch@decagon.com
We need:
  Name/model of instrument
  Quantity
  Symptoms exhibited by the instrument
  Serial number(s) where applicable

## Slide 14: Decagon Tech update list
https://groups.google.com/a/decagon.com/forum/?hl=en#!forum/decagontechupdates